## Slide 1
Rawfile
Calfile
(.CAL)
Rawfile = raw counts file (C:\MTP\Data\NGV\DEEPWAVE\20140705\NG20140705.RAW)
Calfile = calibrations file (C:\MTP\Data\NGV\DEEPWAVE\20140705\NG20140705.CAL)
- flight differences: Noise Diode info
- proj differences: + RollLimit, GEC coefs, Nav temp offset,
frmMTPbin:General /
ReadRAWFile
MTPHfile
(.MTPH)
-MTPYaw,MTPpitch,MTProll (offsets)
-LOSUI array (Sampling Frequencies)
-ElSUI array (Elevation angles)
frmMTPbin:General /
ReadRAWFile
MAM
(Get_MTP_Attitude_Matrix)
CV(1-3,1-14,1-Nscan)
Matrix of raw counts

## Slide 2
CV(1-3,1-14,1-Nscan)
Matrix of raw counts
frmMTPbin:General /
CalculateArrayMAFast
Moving Averager (3 elements)
Replaces value with prior value if
diff between average and raw is
> 500 counts
EV(1-3,1-14,1-Nscan)
Matrix of averaged counts

## Slide 3
Type CFrecord2
' "GENERAL"
Gendate As String * 8 'Analysis Date
WCTdates(1 To 10) As String * 8 'Flight used to generate CAL file
UTstart As Long
UTend As Long
Channels As Integer 'Number of frequency channels
Nel As Integer 'Number of elevation angles
Emissivity As Single 'Emissivity of window
Reflectivity As Single 'Reflectivity of window
DeltaTmin As Single 'Minimum temperature difference between target and sky
' "FIT_INFO"
Nfit As Integer 'Number of fit parameters (including offset)
NP(1 To 5) As String * 6 'Alphanumeric fit parameter (eg Tifa, Ttgt)
GEC(1 To 5, 1 To 5) As Single 'Gain Equation Coefficients (channel, parameter)
GOF(1 To 5) As Single 'Gain Equation Offsets (Channel)
' "WINDOW_CORRECTIONS"
WINcor(1 To 3, 1 To 10) As Single 'Window corrections (Channel, El Angle)
EnableWCT As Boolean
C:\MTP\Data\NGV\{proj}\{fltdate}\NG{fltdate}.BIN file contents (pg 1 of 5)

## Slide 4
' RAW counts editting criteria
CMAcycles As Integer 'Counts smoothing cycles (slow)
CMAcycles2 As Integer 'For Sky Counts only (fast), <= CMAcycles
RFIiterations As Integer 'No of iterations needed to converge (slow)
RFIiterations2 As Integer 'No of iterations needed to converge (fast)
BadCycles As Integer 'No of cycles not satisfying RFI threshold (slow)
Badcycles2 As Integer 'No of cycles not satisfying RFI threshold (fast)
RFIthreshold As Integer 'Maximum counts offset from moving average
MUXthreshold As Single 'Maximum dT in mux temperatures
UseMAforCB As Boolean 'Use Moving Average for Base Counts
UseMAforCS As Boolean 'Use Moving Average for Sky Counts
UseMAforCN As Boolean 'Use Moving Average for Noise Diode Deflection Counts
UseMAforTtgt As Boolean 'Use Moving Average for Target Temperature
UseMAforTifa As Boolean 'Use Moving Average for IF Amp Temperature
' Cycle offsets
aTGToffset(0 To 3) As Integer 'Target Temperature Lead/Lag (+ is lead)
aMXRoffset(0 To 3) As Integer 'Mixer Temperature Lead/Lag
aNDoffset(0 To 3) As Integer 'Noise Diode cycle shift
aNAVoffset(0 To 3) As Single 'OAT cycle shift
aSpare1(0 To 3) As Single
aSpare2(0 To 3) As Single
C:\MTP\Data\NGV\{proj}\{fltdate}\NG{fltdate}.BIN file contents (pg 2 of 5)

## Slide 5
' Gain Thresholds
GeqnMin(1 To 3) As Single 'Gain Equation Min
GeqnMax(1 To 3) As Single 'Gain Equation Max
GnavMin(1 To 3) As Single
GnavMax(1 To 3) As Single
GndMin(1 To 3) As Single
GndMax(1 To 3) As Single
' Channel Weights
ChInfo(1 To 3) As Single
' Fit Region
TBfitX1 As Integer
TBfitX2 As Integer
TBfitY1 As Integer
TBfitY2 As Integer
RHS As Boolean 'TRUE if MTP is on Right Hand Side of a/c
LocHor As Integer 'Scan step of horizon (either 5 or 6)
RAWextension As String * 3 'Extension of file with substituted P/T
Targets As Integer 'Number of reference targets
C:\MTP\Data\NGV\{proj}\{fltdate}\NG{fltdate}.BIN file contents (pg 3 of 5)

## Slide 6
' Noise Diode Temperature Fit
Cnd0(1 To 3) As Single 'Offset
Cnd1(1 To 3) As Single 'First order coefficient
Cnd2(1 To 3) As Single 'Second order coefficient
TrefND As Single 'Reference temperature (Celcius)
' Instrument Attitude
MTPyaw As Single
MTPpitch As Single
MTProll As Single
MTPfiduciary As Single 'Scan Mirror Fiduciary Angle
ElSUI(1 To 10) As Single 'Elevation Angles from SUI file
fEcCount As Integer 'Number of scans with Elevation > fEmax
UseMAforCSgain As Boolean 'Use Moving Average for Sky Counts Based gain
ScanTime As Single 'Average length of a scan from fScanLength(m1, m2)
Expand(1 To 84) As Single 'Room for expansion
C:\MTP\Data\NGV\{proj}\{fltdate}\NG{fltdate}.BIN file contents (pg 4 of 5)

## Slide 7
' FI expansion begins here
NRC As Integer 'Number of RC sets
RCformat(0 To 64) As Integer 'Number indicating RC format (normally = 2 for all sets)
'RCformat = 3, for 10*Nlo observables (ie separate horizon T for each freq)
Reg(0 To 64) As String * 5 '5 character string for each RC set used
RCuse(0 To 64) As Single 'Number of times a give RC has been used in retrieval for a flight
' Len(CFR2)=1744
' Len(REF2)=2000
' Could add 256/11 or 23 more sets of RCs
' Placeholder(0 To 63) As Single 'This makes len(CFR2)=2000
RCformax(65 To 87) As Integer 'Number indicating RC format (normally = 2 for all sets)
'RCformat = 3, for 10*Nlo observables (ie separate horizon T for each freq)
Rex(65 To 87) As String * 5 '5 character string for each RC set used
RCusx(65 To 87) As Single 'Number of times a give RC has been used in retrieval for a flight
Filler As String * 3
End Type
C:\MTP\Data\NGV\{proj}\{fltdate}\NG{fltdate}.BIN file contents (pg 5 of 5)

## Slide 8
13
dT = Tnd - TrefND
For i = 1 To Channels
' Note that the ND deflection (dND) is moving average if UseMAforCN is TRUE
' This is determined in ReadArrayLeadLag
' Cnd0(i) are zero the first time this routine is read
If Cnd0(i) > 0 Then Gnd(i) = dND(i) / (Cnd0(i) * (1# + Cnd1(i) * dT + Cnd2(i) * dT ^ 2))
Next i
14
From BIN file
dND(i) = CV(i, 11, Record) - CV(i, 12, Record)
In REF2readCALFILE: (.BIN file)
Cnd0(i) = CFR2.Cnd0(i)
Cnd1(i) = CFR2.Cnd1(i)
Cnd2(i) = CFR2.Cnd2(i)

## Slide 9
14
Tmix = T `Where i = 4
T = AA + Bb * r(i) + cC * r(i) ^ 2 + DD * r(i) ^ 3
AA = -244.3364635:
Bb = 0.462418:
cC = 0.0000588:
DD = -0.000000013
r(i) = 350# + R7M0 * (Ct(i) - Ct(0)) / DeltaC
DeltaC = Ct(7) - Ct(0)
If DeltaC = 0 Then DeltaC = 12275
R7M0 = 250# 'R(7)-R(0)
Platinum Wire line (raw data):
Pt: 2436 14145 14143 16383 14187 14163 14028 14768
Tnd = T ‘Where i= 6

## Slide 10
Muxs(21) AKA Tmix
15
Because of the Offset & Tifa settings
On the Gain tab, the Geqn calculation
Winds up being a simple Y=Mx+B
Geqn(chan) = FP(2) * GEC(chan, 2) + GEC(chan,1)
FP(2) = Tifa - GOF(2)
Tifa = MV(21, record#)
MV(i, r) = Muxs(i%)
14
From the BIN file
- recalculated/flight

## Slide 11
13
For i = 1 To Channels: g(i) = (g(i) + Gnd(i) * Wtg) / (1 + Wtg): dG(i) = g(i) - Geqn(i): Next i
Wtg = 0.1 'Problem with AltKmTrend if not starting from ground
Unused
' First time thru; give running avg gains: g(1), g(2), g(3) gain equation value
If g(1) = 0 Then g(1) = Geqn(1): g(2) = Geqn(2): g(3) = Geqn(3)
16
From: MTPbin::GainGE:
For i = 1 To Channels
Geqn(i) = GEC(i, 1)
For j = 2 To Nfit
If NP$(j) = "Time" Then ‘ Never Happens
Geqn(i) = Geqn(i) + FP(j) * Exp(-(UV(1, r) - UV(1, 1)) / 1000# / GEC(i, j))
Else
Geqn(i) = Geqn(i) + FP(j) * GEC(i, j)
End If
Next j
Next
15
From BIN file 0 changes/flight!

## Slide 12
RC Configuration: E.g. "c:\MTP\Data\NGV\MPEX\RC\NRCXA156.RCF"
' RC Configuration Information (First record)
Type RCconfiguration
RCformat As Integer
CreationDateTime As Date
RAOBfilename As String * 80
RCfilename As String * 80
RAOBcount As Integer
LR1 As Single 'LR above top of RAOB
zLRb As Single 'LR break altitude
LR2 As Single 'LR above break altitude
RecordStep As Single 'Record Step through available RAOBs
RAOBmin As Single 'Minimum acceptable RAOB altitude
ExcessTamplitude As Single 'Random Excess Noise Level on Ground
Nobs As Integer 'Number of observables
Nret As Integer 'Number of retrieval levels
dZ!(1 To 33) 'Retrieval offset levels wrt flight level
NFL As Integer 'Number of flight levels
Zr!(1 To 20) 'Flight levels (km)
Nlo As Integer 'Number of LO channels
LO!(1 To 3) 'LO frequencies (GHz)
Nel As Integer 'Number of elevation angles
El!(1 To 10) 'Scan mirror elevation angles
Nif As Integer 'Number of IF frequencies
IFoff!(1 To 3, 1 To 16) 'IF frequency offsets (GHz)
IFwt!(1 To 3, 1 To 16) 'Weights assigned to each IF frequency
Spare!(1 To 130)
SURC As String * 4 'SU IFB used to calculate RCs (added 20050128)
CHnLSBloss!(1 To 3) 'CHn LSB RF loss
RAOBbias As Single 'Bias added to RAOB before calculating RCs
CH1LSBloss As Single 'CH1 LSB linear RF loss gradient
' Sensitivity matrix: iRC, NFL, Nlo, Nel
SmatrixN1!(1 To 15, 1 To 3, 1 To 10) 'Linear term
SmatrixN2!(1 To 15, 1 To 3, 1 To 10) 'Quadratic term
End Type
"C:\MTP\Data\NGV\MPEX\RAOB\Templates\ABQ___2013051812.RAOB2 "
"C:\MTP\Data\NGV\MPEX\RAOB\RC\NRCXA156.0000 "
' RC set for one flight level
Type RC_Set_1FL
sBP As Single 'Flight level pressure altitude (hPa)
sOBrms!(1 To 30) '1-sigma apriori observable errors
sOBav!(1 To 30) 'Archive Average observables
sBPrl!(1 To 33) 'Pressure at retrieval levels
sRTav!(1 To 33) 'Average T at retrieval levels
sRMSa!(1 To 33) 'Variance in T at retrieval levels
sRMSe!(1 To 33) 'Formal error in T at retrieval levels
Src!(1 To 33, 1 To 30) '33 retrieval levels, 30 observables
Spare!(1 To 67)
End Type
Records in RCF File: after first
17

## Slide 13
17.5
Function fLevelBelowFL(pALT!)
Dim ib%
ib = 0
Do
ib = ib + 1
Loop Until pALT > FLA(ib) / 100# Or ib >= NFL
fLevelBelowFL = ib
Function fLevelBelowFLwt(pALT!, ib%)
' Calculate weights for above and below flight levels
If FLA(ib - 1) <> FLA(ib) Then
fLevelBelowFLwt = 1 - (pALT * 100# - FLA(ib)) / (FLA(ib - 1) - FLA(ib))
Else
fLevelBelowFLwt = 1
End If
End Function

## Slide 14
template brightness temperatures (and rms & coefs ) are put into a new matrix
- weighted by nearness of aircraft to the two surrounding template flight levels
iBot = fLevelBelowFL(pALT) 'Index of level below FL
iTop = iBot - 1 'Index of level above FL
WtB = fLevelBelowFLwt(pALT, iBot) 'Wt of level below FL
WtT = 1 - WtB 'Wt of level above FL
' Calculate weighted average values of archive average observables
' NRC is number of RC sets
For iNRC = 0 To NRC - 1
iRC = aRC(iNRC)
Call Get_iBot_iTop(iRC, ob(0), iBot, iTop, WtB, WtT)
If iBot = 1 Then
iBot = 2: iTop = 1
End If
For j = 1 To Nobs
OBavgWt(iRC, j) = OBav(iRC, iBot, j) * WtB + OBav(iRC, iTop, j) * WtT
OBrmsWt(j) = OBrms(iBot, j) * WtB + OBrms(iTop, j) * WtT
Coeff1(iRC, j) = SmatrixOB(iRC, iBot, j, 1) * WtB + SmatrixOB(iRC, iTop, j, 1) * WtT
Coeff2(iRC, j) = SmatrixOB(iRC, iBot, j, 2) * WtB + SmatrixOB(iRC, iTop, j, 2) * WtT
Next j
Next iNRC
PALT of aircraft
18
17.5

## Slide 15
Sum1 = 0# ‘(From MTPBin::CalcBiasRmsWrtAA)
Sum2 = 0#
SWi = 0#
NobsP = 0
For j = 1 To Nobs
Wi = 1 / OBrmsWt(j) ^ 2 'Weight Wi
SWi = SWi + Wi 'Sum of Weights
x = (ob(j) - OBavgWt(iRC, j) - OBbias) 'Compare measured TBs to OBavgWt
If Wi > 0# Then
NobsP = NobsP + 1
Sum1 = Sum1 + Wi * x 'Weighted average
Sum2 = Sum2 + Wi * x ^ 2 'Sum of squares
End If
Next j
Xw = Sum1 / SWi 'Calculate Weighted Mean
TBavg(iRC) = Xw
' Correct expression for weighted mean
x = (Sum2 - SWi * Xw ^ 2) / ((NobsP - 1) * SWi / NobsP)
If x >= 0 Then
TBrms(iRC) = Sqr(x) 'Calculate Std Dev about the mean
Else
TBrms(iRC) = TBavg(iRC) 'xxxmjm 2009/02/03
End If
19
Rms values are the same
From template to template?? (thus not indexed by iRC
18

## Slide 16
TAAi(L) = RTav(Index, iBot, L) * WtB + RTav(Index, iTop, L) * WtT
Index = “matched” Template Index - Determined in MTPBin:DetermineRCregion
lnP(iNRC) = 8 * Sqr(TBrms(iRC) ^ 2 + TBavg(iRC) ^ 2) / Nobs 'Sum of ln of Probabilities
Which is the index of the template whose lnP value is the smallest – where lnP
20
19

## Slide 17
21
from: MTPbin:fGetNavT
n = Int(aNAVoffset(C))
If aNAVoffset(C) - n <> 0 Then 'Interpolate
OAT1 = OATV(1, fRecordCheck(Record, n))
OAT2 = OATV(1, fRecordCheck(Record, n + 1))
OAT = OAT1 + (OAT2 - OAT1) * (aNAVoffset(C) - n)
From MTPbin:WriteArray
OATV(1, r) = OATnav
From MTPbin:MapGVtoOther
OATnav = OATnavI
From MTPbin:MapDatV
OATnavI = OATnI + 273.15 'K
From MTPbin:mapDatV
Case (20): OATnI = DatV(i) 'C
Ambient Temperature from IWG1 at
time of scan
From file: MTP\Data\{proj}\{fltdate}\NG{fltdate}.REF
aNAVoffset(0 To 3) As Single 'OAT cycle shift
Always uses the average ~ (-0.14 to -0.18)
- gives Temp 13-15 seconds “ahead” of plane
* at 295 m/s = ~4km
Odd that the individual channels seem to range
From -0.1 to -0.2 meaning that they range in
Distance 4.5km to 4.0km thought that they were
Larger range than that…
Seems like we want to pick a second. And have OAT
On a per channel basis!

## Slide 18
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs ‘(30)
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
33 element array : ob(0) = PALT, ob(1-10 Ch1, 11-20 Ch2, 21-30 Ch3)
MTPbin:MapTAtoOB
For j = 1 To 10: ob(j) = TA(1, j): Next j
For j = 1 To 10: ob(10 + j) = TA(2, j): Next j
For j = 1 To 10: ob(20 + j) = TA(3, j): Next j
MTPbin::TBcalculation
TA(i, j) = OAT + (C(i, j) - C(i, LocHor)) / g(i)
TA(i, j) = TA(i, j) + WINcor(i, j) 'ADD window correction table
22
21
Counts in channel (i) angle (j)
From REF: Determined using spreadsheets late in processing
16

## Slide 19
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
TAAi(L) = RTav(Index, iBot, L) * WtB + RTav(Index, iTop, L) * WtT
From RCF file (MTPio:Rcread) $PROJ\RC\*.RCF
– averaged Temp at retrieval level (sRTav see pg 24)
Index = “matched” Template Index
iBot/iTop = Flight level above/below aircraft (13)
L = Retrieval Level (33)
MTPbin:Get_iBot_iTop(pALT)
Based on Pressure Alt of Aircraft
And PressureAlt of flight levels
- get indices of levels around plane
- calculate weights based on
closeness of plane to flight level
23

## Slide 20
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs ‘(30)
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
RCi(L, j) = rc(Index, iBot, L, j) * WtB + rc(Index, iTop, L, j) * WtT ‘in RetrievalSubs:RetrieveATP
rc!(iRC, n - 1, i, j) = b.Src(i, j) '33 retrieval levels (i), 30 observables (j) : in MTPIO:RCread
' RC set for one flight level
Type RC_Set_1FL
sBP As Single 'Flight level pressure altitude (hPa)
sOBrms!(1 To 30) '1-sigma apriori observable errors
sOBav!(1 To 30) 'Archive Average observables
sBPrl!(1 To 33) 'Pressure at retrieval levels
sRTav!(1 To 33) 'Average T at retrieval levels
sRMSa!(1 To 33) 'Variance in T at retrieval levels
sRMSe!(1 To 33) 'Formal error in T at retrieval levels
Src!(1 To 33, 1 To 30) '33 retrieval levels, 30 observables
Spare!(1 To 67)
End Type
Records in RCF File: after first
24
23
22
18

## Slide 21
Points in temperature profile are
Defined by arrays: (frmATP ATP_Plot)
x = RTp!(j%):
y = RAp!(j%)
Where j% is 1 to Nlev (31 or less)
In Main() RTp(Nlev) = RTavg(i)
* starts where RAavg(i) > 0 (i.e. Nlev != i)
In CalculateAverageProfile():
RTavg(j) = RT(j) + Tcorr
*Tcorr seems to be zero a lot
- only affected by checkbox on TB tab
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
Whole box: 24
25