## Slide 1
Rawfile
Calfile
(.CAL)
Rawfile = raw counts file (C:\MTP\Data\NGV\DEEPWAVE\20140705\NG20140705.RAW)
Calfile = calibrations file (C:\MTP\Data\NGV\DEEPWAVE\20140705\NG20140705.CAL)
- flight differences: Noise Diode info
- proj differences: + RollLimit, GEC coefs, Nav temp offset,
frmMTPbin:General /
ReadRAWFile
MTPHfile
(.MTPH)
-MTPYaw,MTPpitch,MTProll (offsets)
-LOSUI array (Sampling Frequencies)
-ElSUI array (Elevation angles)
frmMTPbin:General /
ReadRAWFile
MAM
(Get_MTP_Attitude_Matrix)
CV(1-3,1-14,1-Nscan)
Matrix of raw counts

## Slide 2
CV(1-3,1-14,1-Nscan)
Matrix of raw counts
frmMTPbin:General /
CalculateArrayMAFast
Moving Averager (3 elements)
Replaces value with prior value if
diff between average and raw is
> 500 counts
EV(1-3,1-14,1-Nscan)
Matrix of averaged counts

## Slide 3: Mux calibrations 1-8 (/1000 * Vscale) -> MV(I,nsamp)
Vscale(8) = -5.1 '-15V PS Vm15
Vscale(7) = 2# 'VCC PS Vp05
Vscale(6) = 5.1 '+15V PS Vp15
Vscale(5) = 7.79 '+24V Syn Vsyn
Vscale(4) = 7.79 '+24V Step Vmtr
Vscale(3) = 2.78 '+8V PS Vp08
Vscale(2) = 1# 'Video V. Vvid
Vscale(1) = -2.73 '-8V PS Vm08

## Slide 4: Mux calibrations (MV) 9-16 (M02 line)
' "Acceler " 'Acc+
' "T Data " 'Tdat
' "T Motor " 'Tmtr
' "T Pod Air " 'Tair
' "T Scan " 'Tsmp
' "T Pwr Sup " 'Tpsp
' "T N/C " 'Tnc
' "T Synth " 'Tsyn

## Slide 5: Mux (MV) calibrations 17-26 (Pt line)
Case 0: R350 = T ‘ 350
Case 1: Ttg1 = T
Case 2: Ttg2 = T
Case 3: Twin = T
Case 4: Tmix = T
Case 5: Tamp = T
Case 6: Tnd = T
Case 7: R600 = T

## Slide 6: MV 25- 40
MV(25, r) = TOMA
MV(26, r) = Zt1
If OATnav > 170 Then
MV(27, r) = OATnav - cTo
Else
MV(27, r) = OATnav
End If
MV(28, r) = OATmms
MV(29, r) = OATmtp
MV(30, r) = pALT
MV(31, r) = gALT
MV(32, r) = rALT
MV(33, r) = Pitch
MV(34, r) = Roll
MV(35, r) = Wspd
MV(36, r) = Wdir (or g(1) set in updateTBrange)
MV(37, r) = VWS (or g(2) set in updateTBrange)
MV(38, r) = GroundSpeed (or g(3) set in updateTBrange)
MV(39, r) = TAS
MV(40, r) = Heading

## Slide 7: Mux calibrations 25- (Pt line)
MV(41, r) = Mach
MV(42, r) = Vzac
MV(43, r) = Tdew
MV(44, r) = Ttotal
MV(45, r) = Pstatic
MV(46, r) = Pdynamic
MV(47, r) = SZA
MV(48, r) = SAAC ‘ Misc tab
MV(49, r) = TTMA
MV(50, r) = TMMA
MV(51, r) = TWMA
MV(52, r) = Zt1
MV(53, r) = LRac
MV(54, r) = LRac2
MV(55, r) = LRac3
MV(56, r) = g(1)
MV(57, r) = g(2)
MV(58, r) = g(3)
MV(59, r) = 0
MV(60, r) = 0
MV(61, r) = 0
MV(62, r) = 0
MV(63, r) = 0
MV(64, r) = 0

## Slide 8: Mux calibrations 25- (Pt line)
' Status tab
MV(65, r) = StatusBits And 128
MV(66, r) = StatusBits And 64
MV(67, r) = StatusBits And 32
MV(68, r) = StatusBits And 16
MV(69, r) = StatusBits And 8
MV(70, r) = StatusBits And 4
MV(71, r) = StatusBits And 2
MV(72, r) = StatusBits And 1

## Slide 10
RC Configuration: E.g. "c:\MTP\Data\NGV\MPEX\RC\NRCXA156.RCF"
' RC Configuration Information (First record)
Type RCconfiguration
RCformat As Integer
CreationDateTime As Date
RAOBfilename As String * 80
RCfilename As String * 80
RAOBcount As Integer
LR1 As Single 'LR above top of RAOB
zLRb As Single 'LR break altitude
LR2 As Single 'LR above break altitude
RecordStep As Single 'Record Step through available RAOBs
RAOBmin As Single 'Minimum acceptable RAOB altitude
ExcessTamplitude As Single 'Random Excess Noise Level on Ground
Nobs As Integer 'Number of observables
Nret As Integer 'Number of retrieval levels
dZ!(1 To 33) 'Retrieval offset levels wrt flight level
NFL As Integer 'Number of flight levels
Zr!(1 To 20) 'Flight levels (km)
Nlo As Integer 'Number of LO channels
LO!(1 To 3) 'LO frequencies (GHz)
Nel As Integer 'Number of elevation angles
El!(1 To 10) 'Scan mirror elevation angles
Nif As Integer 'Number of IF frequencies
IFoff!(1 To 3, 1 To 16) 'IF frequency offsets (GHz)
IFwt!(1 To 3, 1 To 16) 'Weights assigned to each IF frequency
Spare!(1 To 130)
SURC As String * 4 'SU IFB used to calculate RCs (added 20050128)
CHnLSBloss!(1 To 3) 'CHn LSB RF loss
RAOBbias As Single 'Bias added to RAOB before calculating RCs
CH1LSBloss As Single 'CH1 LSB linear RF loss gradient
' Sensitivity matrix: iRC, NFL, Nlo, Nel
SmatrixN1!(1 To 15, 1 To 3, 1 To 10) 'Linear term
SmatrixN2!(1 To 15, 1 To 3, 1 To 10) 'Quadratic term
End Type
"C:\MTP\Data\NGV\MPEX\RAOB\Templates\ABQ___2013051812.RAOB2 "
"C:\MTP\Data\NGV\MPEX\RAOB\RC\NRCXA156.0000 "
' RC set for one flight level
Type RC_Set_1FL
sBP As Single 'Flight level pressure altitude (hPa)
sOBrms!(1 To 30) '1-sigma apriori observable errors
sOBav!(1 To 30) 'Archive Average observables
sBPrl!(1 To 33) 'Pressure at retrieval levels
sRTav!(1 To 33) 'Average T at retrieval levels
sRMSa!(1 To 33) 'Variance in T at retrieval levels
sRMSe!(1 To 33) 'Formal error in T at retrieval levels
Src!(1 To 33, 1 To 30) '33 retrieval levels, 30 observables
Spare!(1 To 67)
End Type
Records in RCF File: after first

## Slide 11
First the template brightness temperatures (and rms & coefs ) are put into a new matrix
- weighted by nearness of aircraft to flight level

## Slide 12
Sum1 = 0# ‘(From MTPBin::CalcBiasRmsWrtAA)
Sum2 = 0#
SWi = 0#
NobsP = 0
For j = 1 To Nobs
Wi = 1 / OBrmsWt(j) ^ 2 'Weight Wi
SWi = SWi + Wi 'Sum of Weights
x = (ob(j) - OBavgWt(iRC, j) - OBbias) 'Compare measured TBs to OBavgWt
If Wi > 0# Then
NobsP = NobsP + 1
Sum1 = Sum1 + Wi * x 'Weighted average
Sum2 = Sum2 + Wi * x ^ 2 'Sum of squares
End If
Next j
Xw = Sum1 / SWi 'Calculate Weighted Mean
TBavg(iRC) = Xw
' Correct expression for weighted mean
x = (Sum2 - SWi * Xw ^ 2) / ((NobsP - 1) * SWi / NobsP)
If x >= 0 Then
TBrms(iRC) = Sqr(x) 'Calculate Std Dev about the mean
Else
TBrms(iRC) = TBavg(iRC) 'xxxmjm 2009/02/03
End If

## Slide 13
TAAi(L) = RTav(Index, iBot, L) * WtB + RTav(Index, iTop, L) * WtT
Index = “matched” Template Index - Determined in MTPBin:DetermineRCregion
lnP(iNRC) = 8 * Sqr(TBrms(iRC) ^ 2 + TBavg(iRC) ^ 2) / Nobs 'Sum of ln of Probabilities
Which is the index of the template whose lnP value is the smallest – where lnP
21

## Slide 14
21
from: MTPbin:fGetNavT
n = Int(aNAVoffset(C))
If aNAVoffset(C) - n <> 0 Then 'Interpolate
OAT1 = OATV(1, fRecordCheck(Record, n))
OAT2 = OATV(1, fRecordCheck(Record, n + 1))
OAT = OAT1 + (OAT2 - OAT1) * (aNAVoffset(C) - n)
From MTPbin:WriteArray
OATV(1, r) = OATnav
From MTPbin:MapGVtoOther
OATnav = OATnavI
From MTPbin:MapDatV
OATnavI = OATnI + 273.15 'K
From MTPbin:mapDatV
Case (20): OATnI = DatV(i) 'C
Ambient Temperature from IWG1 at
time of scan
From file: MTP\Data\{proj}\{fltdate}\NG{fltdate}.REF
aNAVoffset(0 To 3) As Single 'OAT cycle shift
Always uses the average ~ (-0.14 to -0.18)
- gives Temp 13-15 seconds “ahead” of plane
* at 295 m/s = ~4km
Odd that the individual channels seem to range
From -0.1 to -0.2 meaning that they range in
Distance 4.5km to 4.0km thought that they were
Larger range than that…
Seems like we want to pick a second. And have OAT
On a per channel basis!

## Slide 15
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs ‘(30)
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
33 element array : ob(0) = PALT, ob(1-10 Ch1, 11-20 Ch2, 21-30 Ch3)
MTPbin:MapTAtoOB
For j = 1 To 10: ob(j) = TA(1, j): Next j
For j = 1 To 10: ob(10 + j) = TA(2, j): Next j
For j = 1 To 10: ob(20 + j) = TA(3, j): Next j
MTPbin::Tbcalculation
TA(i, j) = OAT + (C(i, j) - C(i, LocHor)) / g(i)
TA(i, j) = TA(i, j) + WINcor(i, j) 'ADD window correction table
22
21
Counts in channel (i) angle (j)

## Slide 16
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
TAAi(L) = RTav(Index, iBot, L) * WtB + RTav(Index, iTop, L) * WtT
From RCF file (MTPio:Rcread) $PROJ\RC\*.RCF
– averaged Temp at retrieval level (sRTav see pg 24)
Index = “matched” Template Index
iBot/iTop = Flight level above/below aircraft (13)
L = Retrieval Level (33)
MTPbin:Get_iBot_iTop(pALT)
Based on Pressure Alt of Aircraft
And PressureAlt of flight levels
- get indices of levels around plane
- calculate weights based on
closeness of plane to flight level
23

## Slide 17
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs ‘(30)
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
RCi(L, j) = rc(Index, iBot, L, j) * WtB + rc(Index, iTop, L, j) * WtT ‘in RetrievalSubs:RetrieveATP
rc!(iRC, n - 1, i, j) = b.Src(i, j) '33 retrieval levels (i), 30 observables (j) : in MTPIO:RCread
' RC set for one flight level
Type RC_Set_1FL
sBP As Single 'Flight level pressure altitude (hPa)
sOBrms!(1 To 30) '1-sigma apriori observable errors
sOBav!(1 To 30) 'Archive Average observables
sBPrl!(1 To 33) 'Pressure at retrieval levels
sRTav!(1 To 33) 'Average T at retrieval levels
sRMSa!(1 To 33) 'Variance in T at retrieval levels
sRMSe!(1 To 33) 'Formal error in T at retrieval levels
Src!(1 To 33, 1 To 30) '33 retrieval levels, 30 observables
Spare!(1 To 67)
End Type
Records in RCF File: after first
24
23
22

## Slide 18
Points in temperature profile are
Defined by arrays: (frmATP ATP_Plot)
x = RTp!(j%):
y = RAp!(j%)
Where j% is 1 to Nlev (31 or less)
In Main() RTp(Nlev) = RTavg(i)
* starts where RAavg(i) > 0 (i.e. Nlev != i)
In CalculateAverageProfile():
RTavg(j) = RT(j) + Tcorr
*Tcorr seems to be zero a lot
- only affected by checkbox on TB tab
In RetrieveATP: ' Do Retrieval using interpolated RCs
For L = 1 To 33
RT(L) = TAAi(L)
For j = 1 To Nobs
RT(L) = RT(L) + RCi(L, j) * (ob(j) - OBavgWt(Index, j)) 'OBavgWt from DetermineRCregion
Next j
Next L
Whole box: 24
25